## Prime L Tech data
Parameter for Prime 30L, 56L
No. |  | Item |  |  | Parameter |  |  |  | Remark / Reference
|  |  | Prime 30L Economic | Prime 30L Standard | Prime 30L Advance | Prime 56L Economic | Prime 56L Standard | Prime 56L Advance
1 |  | Pressure Vessel characteristics
1.01 | 1 | Chamber Capacity | 30 Liters / 7.9 Gallons | 30 Liters / 7.9 Gallons | 30 Liters / 7.9 Gallons | 56 Liters / 14.8 Gallons | 56 Liters / 14.8 Gallons | 56 Liters / 14.8 Gallons | Prime 30 PED data(PED TCF C301-1.1) | Prime 56 PED data(PED TCF C380-1.1)
1.02 | 2 | Chamber Size | Ø 300 x 452 (mm) / Ø 11.8 x 17.8 (inch) | Ø 300 x 452 (mm) / Ø 11.8 x 17.8 (inch) | Ø 300 x 452 (mm) / Ø 11.8 x 17.8 (inch) | Ø 380 x 522 (mm) / Ø 15 x 20.6 (inch) | Ø 380 x 522 (mm) / Ø 15 x 20.6 (inch) | Ø 380 x 522 (mm) / Ø 15 x 20.6 (inch) | Prime 30 PED data(PED TCF C301-1.1) | Prime 56 PED data(PED TCF C380-1.1)
1.03 | 3 | Pressure Vessel Material | Chamber：SUS 316L Stainless Steel , Thinness 2mm Door：SUS 316 Stainless Steel , Thinness 2.5mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Chamber：SUS 316L Stainless Steel , Thinness 2mm Door：SUS 316 Stainless Steel , Thinness 2.5mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Chamber：SUS 316L Stainless Steel , Thinness 2mm Door：SUS 316 Stainless Steel , Thinness 2.5mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Chamber：SUS 316L Stainless Steel , Thinness 2.5mm Door：SUS 316 Stainless Steel , Thinness 3mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Chamber：SUS 316L Stainless Steel , Thinness 2.5mm Door：SUS 316 Stainless Steel , Thinness 3mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Chamber：SUS 316L Stainless Steel , Thinness 2.5mm Door：SUS 316 Stainless Steel , Thinness 3mm Roughness of surface : Polish Gas tungsten arc welding(GTAW) | Prime 30 Chamber drawing | Prime 56 Chamber drawing
1.04 | 4 | Design Presure | 3.0 bar | 3.0 bar | 3.0 bar | 3.0 bar | 3.0 bar | 3.0 bar | PED Cer.
1.05 | 5 | Design Temperature | 144℃ | 144℃ | 144℃ | 144℃ | 144℃ | 144℃ | PED Cer.
1.06 | 6 | Working Presure / Temperature | (-1.0)~ 2.1 bar /0~135 ℃ | (-1.0)~ 2.1 bar /0~135 ℃ | (-1.0)~ 2.1 bar /0~135 ℃ | (-1.0)~ 2.1 bar /0~135 ℃ | (-1.0)~ 2.1 bar /0~135 ℃ | (-1.0)~ 2.1 bar /0~135 ℃ | Prime 30/56 specification(DHF 4-2-1)
1.07 | 7 | Chamber Design Tec. | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Chamber Type：Single chamber w/o jacket Resistant to oxidation, high temperature and high pressure. Gas tungsten arc welding(GTAW) RoHS | Prime 30 Chamber drawing | Prime 56 Chamber drawing
1.08 | 8 | Chamber Door Tec. | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Chamber Door structure：Quick locking cover Concave type Resistant to oxidation | Prime 30 PED data(PED TCF C301-1.1) | Prime 56 PED data(PED TCF C380-1.1)
1.09 | 9 | Insulation system | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Material ：Melamine Foam + AL tape + iron tape thickness 14 mm | Prime 30/56 BOM list(DHF 4-2-5.a)
1.1 | 10 | Design life | 7 years | 7 years | 7 years | 7 years | 7 years | 7 years | Prime 30/56 Life test(DHF 4-8-5)
2 |  | Overall Characteristics
2.01 | 11 | Overall Dimension | 584 (W) x 480 (H) x 670 (D) (mm) 23 (W) x 18.9 (H) x 26.4 (D) (inch) | 584 (W) x 480 (H) x 670 (D) (mm) 23 (W) x 18.9 (H) x 26.4 (D) (inch) | 584 (W) x 480 (H) x 670 (D) (mm) 23 (W) x 18.9 (H) x 26.4 (D) (inch) | 720(W) x 535(H) x 710(D) (mm) | 720(W) x 535(H) x 710(D) (mm) | 720(W) x 535(H) x 710(D) (mm) | Prime 30/56 Installation plan(DHF 4-2-2.c)
2.02 | 12 | Carcass Material | Powder coating Thickness 1mm | Powder coating Thickness 1mm | Powder coating Thickness 1mm | Powder coating Thickness 1mm | Powder coating Thickness 1mm | Powder coating Thickness 1mm | Prime 30 Carcass drawing | Prime 56 Carcass drawing
2.03 | 13 | Preset Program | 15 Programs | 17 Programs | 17 Programs | 15 Programs | 17 Programs | 17 Programs | Prime 30/56 specification(DHF 4-2-1)
2.04 | 14 | Sterilization Temperature | 105~135 C, Adjustable | 105~135 C, Adjustable | 105~135 C, Adjustable | 105~135 C, Adjustable | 105~135 C, Adjustable | 105~135 C, Adjustable | Prime 30/56 specification(DHF 4-2-1)
2.05 | 15 | Temperature Sensor | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | PT100 x 2 pcs, chamber 、folting K type x 3 pcs, steam generator、B-heater、 exhauxt filter (opational) | Prime 30/56 specification(DHF 3-1-1)
2.06 | 16 | Display | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | HMI Digital show temperature, pressure Pressure gauge prossce indicator (green, bule, bule flashing, red) | Prime 30/56 specification(DHF 3-1-1)
2.07 | 17 | Operation Interface | HMI , full color 5' touch panel (screen) | HMI , full color 5' touch panel (screen) | HMI , full color 5' touch panel (screen) | HMI , full color 5' touch panel (screen) | HMI , full color 5' touch panel (screen) | HMI , full color 5' touch panel (screen) | Prime 30/56 specification(DHF 4-2-1)
2.08 | 18 | Altitude Usage | Under 3000M | Under 3000M | Under 3000M | Under 3000M | Under 3000M | Under 3000M | Prime 30/56 specification(DHF 4-2-1)
2.09 | 19 | Net Weight | 68 kg | 73.5 kg | 75 kg | 100 kg | 105.5 kg | 107 kg | Prime 30/56 Packing data(DHF 4-2-3)
2.1 | 20 | Power Supply | 230V ± 10% | 230V ± 10% | 230V ± 10% | 230V ± 10% | 230V ± 10% | 230V ± 10% | Prime 30/56 specification(DHF 4-2-1)
2.11 | 21 | Current | 12.6 A (ampere) | 13.1 A (ampere) | 13.1 A (ampere) | 12.6 A (ampere) | 13.1 A (ampere) | 13.1 A (ampere) | Prime 30/56 specification(DHF 4-2-1)
2.12 | 22 | Fases | 15 A | 15 A | 15 A | 15 A | 15 A | 15 A | Prime 30/56 specification(DHF 4-2-1)
2.13 | 23 | power code | quick socket 16A, European style / USA type | quick socket 16A, European style / USA type | quick socket 16A, European style / USA type | quick socket 16A, European style / USA type | quick socket 16A, European style / USA type | quick socket 16A, European style / USA type | Prime 30/56 specification(DHF 4-2-1)
2.14 | 24 | Power Consumption | Main heater 2000W for steam generator Bend heater 900W | Main heater 2000W for steam generator Bend heater 900W Pipe heater 43W vacuum pump 100W ( standard, advanced) | Main heater 2000W for steam generator Bend heater 900W Pipe heater 43W vacuum pump 100W ( standard, advanced) | Main heater 2000W for steam generator Bend heater 900W | Main heater 2000W for steam generator Bend heater 900W Pipe heater 43W vacuum pump 100W ( standard, advanced) | Main heater 2000W for steam generator Bend heater 900W Pipe heater 43W vacuum pump 100W ( standard, advanced) | Prime 30/56 specification(DHF 4-2-1)
2.15 | 25 | Pre-Vacuum | No | Yes | Yes | No | Yes | Yes | Prime 30/56 specification(DHF 4-2-1)
2.16 | 26 | Post-Vacuum | No | Yes | Yes | No | Yes | Yes | Prime 30/56 specification(DHF 4-2-1)
2.17 | 27 | Pressure & Temperature Control | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 specification(DHF 4-2-1)
2.18 | 28 | Filter | air fiter for lifted vacuum | air fiter for lifted vacuum | air fiter for lifted vacuum Exhaust fiter High presure cooling for Adv. | air fiter for lifted vacuum | air fiter for lifted vacuum | air fiter for lifted vacuum Exhaust fiter High presure cooling for Adv. | Prime 30/56 specification(DHF 4-2-1)
2.19 | 29 | Digital Reorder | Built in memory 2 GB for *** cycle data USB max 64GB | Built in memory 2 GB for *** cycle data Thermal Printer USB max 64GB | Built in memory 2 GB for *** cycle data Thermal Printer USB max 64GB | Built in memory 2 GB for *** cycle data USB max 64GB | Built in memory 2 GB for *** cycle data Thermal Printer USB max 64GB | Built in memory 2 GB for *** cycle data Thermal Printer USB max 64GB | Prime 30/56 specification(DHF 4-2-1)
2.2 | 30 | Connect type | USB to wifi USB to scaner USB to flash memory | RJ-45 for internet USB to wifi USB to scaner USB to flash memory | RJ-45 for internet USB to wifi USB to scaner USB to flash memory | USB to wifi USB to scaner USB to flash memory | RJ-45 for internet USB to wifi USB to scaner USB to flash memory | RJ-45 for internet USB to wifi USB to scaner USB to flash memory | Prime 30/56 specification(DHF 4-2-1)
2.21 | 31 | water trank | 4.2 liter for recycle use | 4.2 liter for recycle use | 4.2 liter for recycle use | 16.1 liter for recycle use | 16.1 liter for recycle use | 16.1 liter for recycle use | Prime 30/56 specification(DHF 3-1-1)
2.22 | 32 | Protection | Overheat Overpressure power Response System Four Door Locking System | Overheat Overpressure power Response System Four Door Locking System Free from Troublesome Agar Clogging | Overheat Overpressure power Response System Four Door Locking System Free from Troublesome Agar Clogging | Overheat Overpressure power Response System Four Door Locking System | Overheat Overpressure power Response System Four Door Locking System Free from Troublesome Agar Clogging | Overheat Overpressure power Response System Four Door Locking System Free from Troublesome Agar Clogging | Prime 30/56 specification(DHF 4-2-1)
2.23 | 33 | Condenser | NA | NA | NA | NA | NA | NA
2.24 | 34 | Type of piping material of the steam and water system | Stainless Steel 304+ Copper+ silicon+PTFE | Stainless Steel 304+ Copper+ silicon+PTFE | Stainless Steel 304+ Copper+ silicon+PTFE | Stainless Steel 304+ Copper+ silicon+PTFE | Stainless Steel 304+ Copper+ silicon+PTFE | Stainless Steel 304+ Copper+ silicon+PTFE | Prime 30/56 specification(DHF 3-1-1)
3 |  | CONTROL SYSTEM for hardware
3.01 | 35 | Control Method | Microcomputer + HMI | Microcomputer + HMI | Microcomputer + HMI | Microcomputer + HMI | Microcomputer + HMI | Microcomputer + HMI | Prime 30/56 specification(DHF 4-2-1)
3.02 | 36 | Control Panel | 5 inch / Touch panel | 5 inch / Touch panel | 5 inch / Touch panel | 5 inch / Touch panel | 5 inch / Touch panel | 5 inch / Touch panel | Prime 30/56 specification(DHF 4-2-1)
3.03 | 37 | connect prot | USB x 3 pcs | USB x 4 pcs RJ-45 for internet | USB x 4 pcs RJ-45 for internet | USB x 3 pcs | USB x 4 pcs RJ-45 for internet | USB x 4 pcs RJ-45 for internet | Prime 30/56 specification(DHF 4-2-1)
3.04 | 38 | Loading Capacity | [6 empty] | Prime 30/56 Operation Manual(DHF 4-2-2.a)
4 |  | CONTROL SYSTEM for Software
4.01 | 39 | Laguage | English | English | English | English | English | English | Prime 30/56 specification(DHF 3-1-1)
4.02 | 40 | Minimum display parameters | Temp:0.1 c Pres:0.001bar | Temp:0.1 c Pres:0.001bar | Temp:0.1 c Pres:0.001bar | Temp:0.1 c Pres:0.001bar | Temp:0.1 c Pres:0.001bar | Temp:0.1 c Pres:0.001bar | Prime 30/56 specification(DHF 3-1-1)
4.03 | 41 | Program | 15 Programs | 17 Programs | 17 Programs | 15 Programs | 17 Programs | 17 Programs | Prime 30/56 specification(DHF 4-2-1) |  | 2.03已有此項目，重複了。
4.04 | 42 | sterilization Programs | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mixed 13. Latex 14. Flash | Prime 30/56 specification(DHF 4-2-1)
4.05 | 43 | test programs | NA | 15. Bowie & Dick Test 16. Leakage | 15. Bowie & Dick Test 16. Leakage | NA | 15. Bowie & Dick Test 16. Leakage | 15. Bowie & Dick Test 16. Leakage | Prime 30/56 specification(DHF 4-2-1)
4.06 | 44 | other program | 17. Cleaning | 17. Cleaning | 17. Cleaning | 17. Cleaning | 17. Cleaning | 17. Cleaning | Prime 30/56 specification(DHF 4-2-1)
4.07 | 45 | Program function | Calculation of F0 value countdown curve chart | Calculation of F0 value countdown curve chart | Calculation of F0 value countdown curve chart | Calculation of F0 value countdown curve chart | Calculation of F0 value countdown curve chart | Calculation of F0 value countdown curve chart | Prime 30/56 specification(DHF 4-2-1)
4.08 | 46 | Functions | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management Free from Troublesome Agar Clogging | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management Free from Troublesome Agar Clogging | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management Free from Troublesome Agar Clogging | ID manergemant ( Adm 1, super uesr 4, nomorl user 15) Calculation of F0 value Calibration ( temperature、pressure、Altitude) Reservation Start Service Remind Day and Time Setting Error code message indication Loads Tracking Management Free from Troublesome Agar Clogging | 規格書須再修改，沒有提到這個欄位內容
4.09 | 47 | Cycle Time | 1. Liquid 1 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 17. Cleaning | 1. Liquid 1 /****/****/***** 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 15. Bowie & Dick Test 16. Leakage 17. Cleaning | 1. Liquid 1 /****/****/***** 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 15. Bowie & Dick Test 16. Leakage 17. Cleaning | 1. Liquid 1 /****/****/***** 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 17. Cleaning | 1. Liquid 1 /****/****/***** 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 15. Bowie & Dick Test 16. Leakage 17. Cleaning | 1. Liquid 1 /****/****/***** 2. Liquid 2 3. Solid 1 4. Solid 2 5. Agar 6. Dissolution 7. user 1 8. User 2 9. Dry only 10. Waste Liquid 11. Waste Solid 12. Waste Mix 13. Latex 14. Flash 15. Bowie & Dick Test 16. Leakage 17. Cleaning | Cycle Time
5 |  | Door Lock SYSTEM
5.01 | 48 | Pressure lock | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | automatic locking control pressure -0.09 bar ~ +0.09 bar Four Door Locking System ~door swich, lock swich…. | Prime 30/56 specification(DHF 3-1-1) |  | 只提到功能，沒提到壓力範圍
5.02 | 49 | Automatic Lock | solenoid lock | solenoid lock | solenoid lock | solenoid lock | solenoid lock | solenoid lock | Prime 30/56 BOM list(DHF 4-2-5.a)
6 |  | HEATING SYSTEM
6.01 | 50 | Sterilization heating | Steam generator | Steam generator | Steam generator | Steam generator | Steam generator | Steam generator | Prime 30/56 specification(DHF 3-1-1)
6.02 | 51 | Power | 2900 w | 2900 w | 2900 w | 2900 w | 2900 w | 2900 w | Prime 30/56 specification(DHF 4-2-1)
6.03 | 52 | Bend heater | 900 w | 900 w | 900 w | 900 w | 900 w | 900 w | Prime 30/56 specification(DHF 4-2-1)
7 |  | Pipe and water system
7.01 | 53 | Type of piping material of the steam | Stainless Steel 304+ Copper+PTFE | Stainless Steel 304+ Copper+PTFE | Stainless Steel 304+ Copper+PTFE | Stainless Steel 304+ Copper+PTFE | Stainless Steel 304+ Copper+PTFE | Stainless Steel 304+ Copper+PTFE | Prime 30/56 specification(DHF 3-1-1)
7.04 | 56 | water trank | 4.2 liter, PP material water level high and low of clean water tank | 4.2 liter, PP material water level high and low of clean water tank | 4.2 liter, PP material water level high and low of clean water tank | 16.1 liter,Stainless Steel 304 material water level high and low of clean water tank | 16.1 liter,Stainless Steel 304 material water level high and low of clean water tank | 16.1 liter,Stainless Steel 304 material water level high and low of clean water tank | Prime 30/56 specification(DHF 3-1-1)
7.05 | 57 | fill water | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | automatic fill water for clean water tank, water presure 2-7 bar automatic fill water for steam generator by fill water motor | Prime 30/56 specification(DHF 3-1-1)
7.06 | 58 | Drain | quick drain port | quick drain port | quick drain port | quick drain port | quick drain port | quick drain port | Prime 30/56 specification(DHF 3-1-1)
8 |  | Air remove / dry / insulation
8.01 | 59 | gravity (no vacuum pump type) | gravity pulse | gravity pulse | gravity pulse | gravity pulse | gravity pulse | gravity pulse | Prime 30/56 Operation Manual(DHF 4-2-2.a)
8.02 | 60 | Pre Vaccum | NA | vacuum pump spec. Max Flow XX LPM, MAX Vacuum -0.XX bar Pulse 2 times Max. vacuum value -0.9 bar | vacuum pump spec. Max Flow XX LPM, MAX Vacuum -0.XX bar Pulse 2 times Max. vacuum value -0.9 bar | NA | vacuum pump spec. Max Flow XX LPM, MAX Vacuum -0.XX bar Pulse 2 times Max. vacuum value -0.9 bar | vacuum pump spec. Max Flow XX LPM, MAX Vacuum -0.XX bar Pulse 2 times Max. vacuum value -0.9 bar | Prime 30/56 specification(DHF 3-1-1)
8.03 | 61 | Post Vacuum | NA | Yes | Yes | NA | Yes | Yes | Prime 30/56 specification(DHF 3-1-1)
8.04 | 62 | Heater | Bend-Heater 900W | Bend-Heater 900W | Bend-Heater 900W | Bend-Heater 900W | Bend-Heater 900W | Bend-Heater 900W | Prime 30/56 specification(DHF 4-2-1)
8.05 | 63 | Insulation | keep warm 0h~24h Temperature 45-60°C | keep warm 0h~24h Temperature 45-60°C | keep warm 0h~24h Temperature 45-60°C | keep warm 0h~24h Temperature 45-60°C | keep warm 0h~24h Temperature 45-60°C | keep warm 0h~24h Temperature 45-60°C | Prime 30/56 Operation Manual(DHF 4-2-2.a)
9 |  | AIR FILTER
9.01 | 64 | Vacuum remove air filter | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | Prime 30/56 specification(DHF 4-2-1)
9.02 | 65 | Exhaust Filter | NA | NA | filter chamber material SUS304 sterilization 150 cycle times @ 134C temperature sensor K type | NA | NA | filter chamber material SUS304 sterilization 150 cycle times @ 134C temperature sensor K type | Prime 30/56 specification(DHF 4-2-1)
9.03 | 66 | High pressure cooling air filter | NA | NA | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | NA | NA | HEPA air filter 0.03μm EN 13060, BFE and VFE filtration efficiency as high as 99.999% | Prime 30/56 specification(DHF 4-2-1)
10 |  | COOLING SYSTEM
10.01 | 67 | Exhaust | Adjustable Level 0 - 10 | Adjustable Level 0 - 10 | Adjustable Level 0 - 10 | Adjustable Level 0 - 10 | Adjustable Level 0 - 10 | Adjustable Level 0 - 10 | Prime 30/56 specification(DHF 4-2-1)
10.02 | 68 | Fan cooling | DC 24V x 1 pcs | DC 24V x 3 pcs | DC 24V x 3 pcs | DC 24V x 1 pcs | DC 24V x 3 pcs | DC 24V x 3 pcs | Prime 30/56 specification(DHF 4-2-1)
10.03 | 69 | Condenser | NA | NA | NA | NA | NA | NA
10.04 | 70 | High pressure cooling | NA | NA | pressure adjustment valve (with air-water separation), @ 2.0 bar piping cooling rate | NA | NA | pressure adjustment valve (with air-water separation), @ 2.0 bar piping cooling rate | Prime 30/56 specification(DHF 4-2-1)
11 |  | Data Recorder & Traceable System
11.01 | 71 | Printer | NA | Integrated thermal printer | Integrated thermal printer | NA | Integrated thermal printer | Integrated thermal printer | Prime 30/56 specification(DHF 4-2-1)
11.02 | 72 | Printer Paper | NA | Thermal paper x 5 pcs | Thermal paper x 5 pcs | NA | Thermal paper x 5 pcs | Thermal paper x 5 pcs | Prime 30/56 specification(DHF 4-2-1)
11.03 | 73 | Memory | USB memory max. 64GB | USB memory max. 64GB | USB memory max. 64GB | USB memory max. 64GB | USB memory max. 64GB | USB memory max. 64GB | Prime 30/56 specification(DHF 4-2-1)
11.04 | 74 | Recorder reading software | FileChecker | FileChecker | FileChecker | FileChecker | FileChecker | FileChecker | Prime 30/56 Operation Manual(DHF 4-2-2.a)
11.05 | 75 | Traceable mode | Pre-scan and post scan scaner bar code printer (@150℃) | Pre-scan and post scan scaner bar code printer (@150℃) | Pre-scan and post scan scaner bar code printer (@150℃) | Pre-scan and post scan scaner bar code printer (@150℃) | Pre-scan and post scan scaner bar code printer (@150℃) | Pre-scan and post scan scaner bar code printer (@150℃) | Prime 30/56 specification(DHF 3-1-1)
11.06 | 76 | IoT server | IoT connect by wifi | IoT connect by wifi or RJ-45 | IoT connect by wifi or RJ-45 | IoT connect by wifi | IoT connect by wifi or RJ-45 | IoT connect by wifi or RJ-45 | Prime 30/56 specification(DHF 4-2-1)
12 |  | Accessories
12.01 | 77 | basic panel | Standard | Standard | Standard | Standard | Standard | Standard | Prime 30/56 specification(DHF 4-2-1)
12.02 | 78 | software | Standard | Standard | Standard | Standard | Standard | Standard | Prime 30/56 specification(DHF 4-2-1)
12.03 | 79 | Wire-Mesh Basket | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.04 | 80 | Waste Sterilization Box | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.05 | 81 | Tray Set | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.06 | 82 | Barcode Scanner | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.07 | 83 | Barcode Printer | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.08 | 84 | Water Distiller | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.09 | 85 | RO Water Filter | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.1 | 86 | Wi-Fi Dongle | Optional | Optional | Optional | Optional | Optional | Optional | Prime 30/56 specification(DHF 4-2-1)
12.11 | 87 | SteriProcess IoT Server | Standard / Optional | Standard / Optional | Standard / Optional | Standard / Optional | Standard / Optional | Standard / Optional | Prime 30/56 specification(DHF 4-2-1)
13 |  | Documents:
User manual
13.01 | 88 | Catalog with Technical Sheet | Yes | Yes | Yes | Yes | Yes | Yes
13.02 | 89 | Operation Manual | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 Operation Manual(DHF 4-2-2.a)
13.03 | 90 | Service Manual | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 Service Manual(DHF 4-2-2.b)
13.04 | 91 | Warranty of manufacture defect (since installation) | one Year | one Year | one Year | one Year | one Year | one Year
13.05 | 92 | Spare part (accessories) Manual | Yes | Yes | Yes | Yes | Yes | Yes
13.06 | 93 | Availability of providing spare parts for maintenance | Yes | Yes | Yes | Yes | Yes | Yes
13.07 | 94 | Provide product operation training, preventive maintenance, frequently defect solution to the operator | Yes | Yes | Yes | Yes | Yes | Yes
14 |  | Standards and Directives
14.01 | 95 | ISO 13485:2016 Quality Management System - Medical Devices | Yes | Yes | Yes | Yes | Yes | Yes | ISO 13485 Cer.
14.02 | 96 | PED 2014/68/EU Pressure Equipment Directive | Yes | Yes | Yes | Yes | Yes | Yes | PED Cer.
14.03 | 97 | ASME BPVC Section VIII-1 Rules for Construction of Pressure Vessels | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 specification(DHF 4-2-1)
14.04 | 98 | EN/IEC 61010-1 Safety Requirements for measurement control and laboratory use | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 LVD Cer.(DHF 4-8-2)
14.05 | 99 | EN/IEC 61010-2-040 Safety requirements for sterilizers used to treat medical materials | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 LVD Cer.(DHF 4-8-2)
14.06 | 100 | EN/IEC 61326-1 EMC | Yes | Yes | Yes | Yes | Yes | Yes | Prime 30/56 EMC Cer.(DHF 4-8-1)
14.07 | 101 | RoHS Restriction of Hazardous Substances Directive 2011/65/EU | Yes | Yes | Yes | Yes | Yes | Yes | N/A
Note on D5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on E5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on F5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on G5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on H5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on I5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已與RD確認
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD確認
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD確認
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD確認
Note on D7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on E7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on F7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on G7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on H7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on I7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 3/1 已確認
Note on D8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on E8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on F8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on G8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on H8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on I8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已跟小D確認2.8 因安規報告的關係
Note on D9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on E9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on F9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on G9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on H9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on I9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 小D回覆 正確是142
Note on D10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on E10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on F10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on G10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on H10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on I10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 已與RD 確認規格
Note on C24: Jiasheng(黃嘉聖): 重量再確認
Note on C44: Jiasheng(黃嘉聖): 再確認負載數量

## Cycle Time
Prime 30L Cycle Time
Program | Loading |  | Condition | Time
Solid | N/A | 121 | S : 20 min、Dry 20 min | 2:27
|  | 134 | S : 20 min、Dry 20 min | 1:57
Agar | 250ml*3 | 121 | S : 30 min、EX10、Ins.60 、1H、CT80、FT95 | 2:42
Dissolution | 250ml*3 | 80 | S : 15 min、CT80、FT95 | 0:36
B&D |  |  |  | 2:06
Flash | 7KG |  |  | 0:41
Latex | 7KG |  |  | 1:31
Liquid | 250ml*3 | 121 | S : 30min、EX10、CT80、FT95 | 1:44
Liquid | 250ml*3 | 134 | S：15min、EX0、CT80、FT95 | 1:20
Waste Solid | 7KG |  |  | 1:49
Waste Mixed | 250ml*1+1KG |  |  | 1:27
Waste Liquid |  |  |  | 1:35
USER | 7KG |  | S：10min、EX0、CT80、FT95 | 1:11